Calibration Worksheet

For

Pressure Measuring Devices

Date:  _______________ (calibration frequency not to exceed 90 calendar days for leak decay testing)

Equipment:  	              	Digital Manometer*         	           Mechanical Gauge**

Serial Number: __________________________ (serial number to be permanently engraved in device)

20% of full scale: ____________(incline)  ____________ (actual reading gauge being calibrated)

50% of full scale: ____________(incline)  ____________  (actual reading gauge being calibrated)

80% of full scale: ____________(incline)  ____________  (actual reading gauge being calibrated)

Calibrated gauge must be accurate within 2% of reference gauge (incline).  Scale of incline manometer must be the same as the gauge being calibrated.

*Digital manometers with a range of 0-10 inches must have an accuracy of .5% and digital manometers with a range of 0-20 inches must have an accuracy of .25%.  Please refer to manufacturers specifications for gauge accuracy rating.

**Mechanical gauges with pressure ranges of 0-2.0” wc, 0-1.0” wc, and 0-0.50” wc shall have a minimum 4” face, shall be accurate to 3% of full scale, and the maximum incremental graduations shall be .05”wc.

Company Name: __________________________________________________

Person conducting calibration: _______________________________________

Signature:  _______________________________________________________

NOTE:  Only original, unaltered forms will be accepted (no white out corrections or crossed out information in any section will be allowed).